## Slide 1
Phase contrast
Jerome Mertz

## Slide 2
time
frequency
F.T.
It all starts from this….
position
(along x)
momentum
(along x)
F.T.

## Slide 3
Fourier transform relationships (diffraction)
Dx Dqx ~ l/n
Dx
Dqx
small
large
l
Dqx
Dx
large
small

## Slide 4
Dx Dqx ~ l/n
Fourier transform relationships (diffraction)
Dqx
Dx
l
Unscattered light
Scattered light
Dx

## Slide 5
Sample
Resolution
Imaging
device
Actual sizes

## Slide 6
Imaging
device
Resolution
Sample
Actual sizes

## Slide 7
Imaging
device
Resolution
Sample
Apparent sizes
Minimum apparent size
(a.k.a. Resolution)

## Slide 8
Imaging
device
Resolution
1/2
Apparent sizes
Minimum apparent size
(aka. Resolution)
Sample
1/2
1/2
Numerical aperture 
1/2
Air:
Water:
Glass:

## Slide 9
Absorption
Scattering
Back to phase….

## Slide 10
Zernike phase contrast
Nomarski differential interference contrast (DIC)
Oblique field microscopy
Tricks to convert phase to intensity

## Slide 11
n = index of refraction
wave-number

## Slide 12
n = index of refraction
wave-number

## Slide 13
phase shift
wave-number

## Slide 14
Imaging system
Intensity:
 small
What if ballistic field is phase-shifted relative to scattered field?
 not so small
CCD
Object
ballistic
scattered
Field:
Born approximation

## Slide 15
Imaging system
Rule of thumb:
(i.e. need high NA to see small structures)
CCD
Object
Dx Dqx ~ l/n
Dqx

## Slide 16
spatial frequency (sample)
wave-number (light)
(i.e. need high NA to see high spatial frequencies)

## Slide 17
Scattered field
Ballistic field

## Slide 18
Object plane
Fourier plane
Image plane
(conjugate to object plane)
Ballistic
Scattered
maps angle to position

## Slide 19
Object plane
Fourier plane
Image plane
(conjugate to object plane)
Ballistic light here

## Slide 20
Object plane
Fourier plane
Image plane
(conjugate to object plane)
Zernike phase contrast
1) Add /2 phase shift to ballistic light in Fourier plane:
Ballistic light here

## Slide 21
Object plane
Fourier plane
Image plane
(conjugate to object plane)
Zernike phase contrast
1) Add /2 phase shift to ballistic light in Fourier plane:
2) Attenuate ballistic light to improve contrast:
similar in magnitude
Ballistic light here

## Slide 22
Object
Illumination must be spatially coherent
Köhler illumination doesn’t work!
Problem !
Phase plate

## Slide 23
Incoherent (spatially random phases)
This solution is not ideal because illumination power is weak
Köhler illumination
Object
Phase plate
Pinhole
Spatially coherent
Solution:

## Slide 24
Annular phase-plate
Annular illumination
Alternative solution:
Spatially coherent (sort of)
Advantages of annular illumination:
More illumination throughput
Fewer backreflections from optics
Better resolution (though difficult to interpret)
Object

## Slide 25
Annular phase-plate
Annular illumination
Alternative solution:
Phase contrast condenser
Phase contrast objective
Spatially coherent (sort of)
Advantages of annular illumination:
More illumination throughput
Fewer backreflections from optics
Better resolution (though difficult to interpret)

## Slide 26
Brightfield
Zernike PC
Nomarski DIC

## Slide 27
Zernike phase contrast
Nomarski differential interference contrast (DIC)
Oblique field microscopy
Tricks to convert phase to intensity

## Slide 28
Cross-polarized
Intensities add
Random phases
Intensities add
Well-defined phases
Fields add
More about phase…

## Slide 29
Polarizer
Polarizing beamsplitter
Wollaston prism
Nomarski prism
mrad
…and polarization

## Slide 30
=
=
=
or
In phase, equal amplitude
Linear polarization 45o
In phase, unequal amplitude
Linear polarization
Out of phase, equal amplitude
Circular polarization
Elliptic polarization

## Slide 31
CCD
Nomarski
Spatially coherent illumination
(polarization 45o)
Transparent sample

## Slide 32
Spatially coherent illumination
(polarization 45o)
CCD
Nomarski
Polarizer -45o
(analyzer)
Transparent sample

## Slide 33
Spatially coherent illumination
(polarization 45o)
CCD
Nomarski
Polarizer -45o
(analyzer)
Sensitive to phase gradient
Problem: how to get bright coherent illumination with a lamp?
phase shifted
Key condition:
Scattering sample

## Slide 34
CCD
Nomarski
Polarizer -45o
(analyzer)
Polarizer +45o
Polarized Köhler illumination?
lamp
Scattering sample

## Slide 35
CCD
Nomarski
Polarizer -45o
(analyzer)
Polarizer +45o
Polarized Köhler illumination?
lamp
Scattering sample
Doesn’t work because Köhler illumination not spatially coherent
Random phases:

## Slide 36
CCD
Nomarski
Polarizer -45o
(analyzer)
Polarizer +45o
lamp
Nomarski
Trick: DIC microscopy
Polarized Köhler illumination with Nomarski prism provides bright uniform illumination that is locally pairwise coherent
Pairwise correlated:
Scattering sample

## Slide 37
Köhler illumination
(spatially incoherent and unpolarized)
Polarizer 45o
+
DIC illumination

## Slide 38
Polarizer 45o
+
DIC illumination
Pairwise coherent
+
Nomarski prism
Köhler illumination
(spatially incoherent and unpolarized)

## Slide 39
CCD
Nomarski
Polarizer -45o
(analyzer)
Polarizer +45o
lamp
Nomarski
Trick: DIC microscopy
Polarized Köhler illumination with Nomarski prism provides bright uniform illumination that is locally pairwise coherent
Pairwise correlated:
Scattering sample

## Slide 40
Brightfield
Zernike PC
Nomarski DIC

## Slide 41
Very useful in neuroscience

## Slide 42
Zernike phase contrast
Nomarski differential interference contrast (DIC)
Oblique field microscopy
Tricks to convert phase to intensity

## Slide 43
Transmission
geometry
Thin slice
Phase contrast
But there’s a problem….

## Slide 44
Reflection
geometry
Thick tissue
Phase contrast
?

## Slide 45
To camera
LED
LED
Thick mouse brain slab
(depth 50µm)
Oblique back-illumination microscopy (OBM)

## Slide 46
Camera
LED
LED
Tissue
OBM microendoscopy
70µm working distance
2 shots per frame

## Slide 47
80 µm
30,000 fibers
Mauna-Kea Technology
3mm
OBM with microendoscope (chick embryo)

## Slide 48
30m
Frame rate: 25 Hz
30,000 fibers
Mauna-Kea Technology
3mm
OBM with microendoscope (chick embryo)

## Slide 49
Mouse skin tissue
(microscope configuration)

## Slide 50
Cow trachea

## Slide 51
How does it work?
Aperture
Oblique illumination

## Slide 52
Aperture
Oblique illumination
How does it work?

## Slide 53
Aperture
Oblique illumination
How does it work?

## Slide 54
Aperture
Oblique illumination
How does it work?

## Slide 55
LEFT
45m bead
Scattering phantom:
2m beads mixed in
1cm thick agarose gel
ls = 75m ; ls*  1mm

## Slide 56
LEFT
SUM
DIFFERENCE
RIGHT
45m bead
Scattering phantom:
2m beads mixed in
1cm thick agarose gel
ls = 75m ; ls*  1mm

## Slide 57
En-face OCT
OBM
Comparison OCT and OBM
(Mouse flank skin)
Image acquired by Light-CT Scanner,
LLTech (Paris)

## Slide 58
an aside…

## Slide 59
Widefield microscope
Source
Array detector
Principle of reciprocity
Large area detector
Scanning source
Scanning microscope

## Slide 60
Widefield microscope
Source
Array detector
Large area detector
Scanning source
Principle of reciprocity
Scanning microscope

## Slide 61
Widefield microscope
Scanning microscope
Source
Array detector
Large area detector
Scanning source
Principle of reciprocity

## Slide 62
To camera
LED
LED
OBM
If this works…
|
+
Phase gradient
Amplitude
From laser
Scanning OBM
equivalent
Principle of reciprocity
then this must also work

## Slide 63
Davison Lab, Boston University
TPEF
sOBM
GFP-labeled mouse cortex
Scanning OBM with TPEF
FOV = 80m
depth 

## Slide 64
CARS
sOBM
Coté Lab, University Laval
Rat sciatic nerve
Scanning OBM with CARS
FOV = 80m

## Slide 65
Endomicroscopy (Scanning confocal microscopy)
Fluorescence
Mouse colon epithelium
…in collaboration with Ritt Lab
30,000 fibers
2.8mm
Pinhole

## Slide 66
Endomicroscopy (Fluorescence + Phase)
Fluorescence
Phase
Mouse colon epithelium
…in collaboration with Ritt Lab
Combined
30,000 fibers
2.8mm
Reflecting
pinhole